## Slide 1
TI-ADC Calibration on ZCU102
From Hardware Setup to a Closed-Loop Calibration Pipeline
Group Meeting — August 2026
Notes: Opening: Last time I showed that the ADC and DAC hardware were set up. Today I will show how we turned that setup into a complete automatic calibration loop running on the board.

## Slide 2
Where We Were Last Time
• AD9695 dual ADC: JESD204B link working
• AD9164 DAC + DPG: 2.6 GSPS waveform loading
• Raw ADC capture working
• UDP transfer from board to PC working
• But: no automatic calibration loop on hardware
Status: Data acquisition path working Calibration loop: NOT YET CLOSED
Notes: At the end of last meeting, the bench was working as a data acquisition system, but calibration was still not closed-loop.

## Slide 3
System Architecture
DPG → AD9164 DAC → analog front end → AD9695 ADC → JESD204B → ZCU102 PS
Reference input: 199.375 MHz tone + balanced dither pulses
• ADC: 1300 MSPS, two parallel channels A and B
• DAC: 2600 MSPS, fs_dac / fs_adc = 2
• Data path: JESD204B → AXI DMA → PS memory → UART / UDP export
• Correction path: software offset/gain + AD9695 fine clock delay for skew
Notes: The idea is simple: we inject a known reference waveform into both channels and use the measured mismatch to drive calibration.

## Slide 4
Five-Stage Calibration Pipeline
Timing
align captures to uploaded reference
→
Offset
software negative feedback
→
Gain
software negative feedback
→
Skew
actuator characterization + closed-loop correction
→
Performance
SNDR / SFDR / ENOB reporting
Board command: adc -cal → runs all five stages automatically on the ZCU102
Notes: Everything runs inside the ZCU102 firmware. The PC is used only for uploading the reference and exporting results.

## Slide 5
Stages 1–3: Converged on the Board
Stage | Result | Key number
Timing | PASS | correlation 0.99902, 10/10 frames
Offset | CONVERGED | verification residual +0.59 code
Gain | CONVERGED | verification error −0.00054
These stages confirm that the reference path, acquisition path, and software correction model work on the real hardware.
Notes: These three stages confirm that the reference path, acquisition path, and the software correction model are all working on the real hardware.

## Slide 6
Stage 4 — Skew Was the Hard Part
Actuator step is comparable to measurement noise
• Initial B−A skew: about −60 ps
• Required tolerance: ±7.69 ps
• Actuator step: a few ps per register code
• Batch noise: 7–17 ps
• Old approach: fixed ±1-code probe
• Result: response could not be separated from noise
• Calibration failed at actuator characterization
Old failure: ACTUATOR CHARACTERIZATION FAILED response not distinguishable from noise
Notes: Unlike offset and gain, skew correction is limited by how small one actuator step is. A single-code probe was below the measurement noise floor.

## Slide 7
Stage 4 — Adaptive Characterization Fixed Skew
Fix
• Re-measure fresh baseline before every probe
• Escalate probe amplitude: 1 → 2 → 4 → 8 codes
• Keep tone estimator and controller unchanged
Final result
• Skew: −60.04 ps → −5.96 ps
• Register: 24 → 32 → 33 → 34
• 2/2 consecutive passes
• Measured step: 3.96 ps/code
Overall calibration passed Stage 4 and continued to Performance
Notes: The final 8-code probe was large enough to be measured reliably, and then the normal controller closed the loop.

## Slide 8
Stage 5 — Correct Channel Polarity
Stage-4 INVERTED relation propagated into performance analysis
Metric | Before | After
A/B correlation | −0.99995 | +0.99992
A/B RMSE | 768 codes | 18.2 codes
Parallel average SNDR | 3.02 dB | 37.8 dB
ENOB | 0.21 bits | ~6.0 bits
After the fix, the combined output no longer collapses.
Notes: After propagating the Stage-4 polarity decision into Stage 5, the combined output no longer collapses.

## Slide 9
Limitations and Next Steps
Limitations
• Dither fine-skew remains INVALID
• Edge disagreement ~168 ps vs 23.1 ps gate
• Analog chain disperses the injected pulse
• A/B gain mismatch: −1.37%
• Parallel average gives no +3 dB
Next steps
• Re-run board with fixed firmware for a clean full log
• Test 199.375 MHz linear-triangle dither waveform
• Add pulse-width diagnostics to dither estimator
• Investigate per-channel gain correction
• Study JESD re-lock phase jumps
Notes: These are quantified limitations, not hidden failures. The main calibration path is working.

## Slide 10
Conclusion
• Hardware setup is now a full on-board calibration loop
• Timing, offset, gain, skew, and performance all PASS
• Stage-4 skew closed-loop correction is the main new result
• Known limitations are documented and measurable
• Next: dither fine-skew and per-channel gain improvement
Thank you. Questions?
Notes: Thank you. I am happy to take questions.